Colorimetry

Colorimetry is a photometric technique. Photometry is the most common analytical technique used in clinical biochemistry. The principle of photometry is based on the physical laws of radiant energy or light. The photometric principle asserts that when an incident light beam of intensity, Io passes through a solution, some of the light is reflected (Ir), some is absorbed, and the rest of the light is transmitted (It).

Io=Ir + Ia + It

Photometric principles are applied in several kinds of analytical measurements.

1. Measurement of absorbed or transmitted light: Colorimetry, spectrophotometry, atomic absorption, turbidometry.

2. Measurement of emitted light, Flame emission photometry: Fluorometry.

Colorimetry is a scientific technique used to determine the concentration of colored compounds in solutions. Many methods for quantitative analysis of biological materials are based upon the production of a colored compounds in solution, the intensity of which is used as a measure of concentration.

The colorimeter is based on Beer-Lambert’s law which states that the absorption of light passed through a medium is proportional to its concentration. Let’s discuss these laws;

Beer’s Law

According to Beer’s law, the amount of light absorbed is directly proportional to the concentration of the solute in the solution. Or when a ray of monochromatic light passes through a solution, the intensity of light transmitted decreases exponentially as the concentration of the medium increases.

I = I0 e–K1 c where,

I = intensity of incident light

Io = intensity of transmitted light

C= concentration of the solution

Lambert’s Law

According to Lambert, the amount of light absorbed is directly proportional to the length and thickness of the solution under consideration Or when monochromatic light passes through a solution, the intensity of light transmitted decreases exponentially with increasing path length.

I = I0 e–K2 l

I = intensity of incident light

Io = intensity of transmitted light

C= concentration of the solution

l= length/thickness of the solution

The ratio of intensities is known as the transmittance (T) and it is usually expressed as a percentage.

Percent transmittance = I/I0 × 100 = e–K3 cl

This is not very convenient which is evident from the negative exponential curve obtained by plotting percent transmittance against concentration.

If logarithms are taken instead of a ratio, then loge I/I0 = K3 c l

Log10 I/I0 = K3 c l

The expression Log10 I/I0 is known as extinction (E) or absorbance (A).

The colorimeter works on the basis of combination of the two laws, the Beer-Lambert’s law which states that the amount of light absorbed by a color solution is directly proportional to its concentration and the length of a light path through it. Mathematically represented as;

Aα cl

A=εcl   where

A = absorbance of the solution

ε= molar absorptivity

c= concentration of the solution

l= distance travelled by light in the solution

Filters and wavelength

A colorimeter consists of changeable optics filters which select the wavelength that the solution absorbs the most in order to maximize accuracy. A colorimeter uses the wavelength variation ranging from 400 to 700nm.

In colorimetric determinations, specific reagents are used which react with the specific component to form a colored complex. The concentration of the colored complex is directly proportional to the component in the specimen.

Components of a colorimeter

- Light Source: The source of light should produce energy with enough intensity to cover the entire visible spectrum (380-780 nm). Commonly, Tungsten lamps are used as a light source for measurement in the visible spectrum and near-infrared ranges. Halogen deuterium is suitable for measurement in the UV range (200-900 nm).

- Condensing lens: Parallel beam of light emerges from condensing lens after the light passes through slit incidents on it.

- Monochromator: It filters the monochromatic light from polychromatic light, which absorbs unwanted light wavelengths and permits only monochromatic light. There are of three types: prism, grating, and glass.

  - Prism: It facilitates the refraction of light when it passes from one medium to another.

  - Glass: It selectively transmits light in certain ranges of wavelengths.

  - Gratings: These are made of graphite, which separates light in different wavelengths.

- Cuvette (Sample cell): The monochromatic light from the filter passes through the colored sample solution placed in the cuvette.  Their sizes

range from square, and rectangle to round and have a fixed diameter of 1cm. There are of three types based on the substances these are made of: Glass, Quartz, and Plastic cuvette.

  - Glass cuvettes are cheap and absorb light of 340 nm wavelength.

  - Quartz cuvettes facilitate entry of both lights of UV and visible ranges.

  - Plastic cuvettes are cheaper, easily scratched, and have shorter lifespans.

- Photocell (Photodetector): These photosensitive devices measure light intensity by converting light energy into electrical energy.

- Galvanometer: The electrical signal generated in a photocell is detected and measured by a galvanometer. It displays optical density (OD) and percentage transmission.

Procedure of colorimetric estimation of concentration

In colorimetric estimation, it is necessary to prepare 3 solutions; the blank usually water or reagent to eliminate the effect of light absorption, the standard which is a solution with known concentration and the test with unknown concentration.

Following a combination of the Beer and lamberts law, the concentration of the test solution can be determined.

A standard curve is then plotted to see if the method of assaying follows Beer’s law i.e., whether the absorbance increases in a linear manner with increase in concentration.

The standard curve is a plot of y for optical density against x for the concentration of the standard solution. The concentration of the unknown can be obtained by extrapolating the known optical density or absorbance in the graph.

Example

To prepare a 1% standard solution of glucose, i.e. 1g/dl

You first have to make different dilutions of the standard solution using the general formula for obtaining concentration C1xV1=C2xV2

After preparing these concentrations of Std, their O.Ds or absorbances are measured. A standard curve can be prepared as shown below.

Criteria for getting satisfactory colorimetric estimations

- Stability of the color: Color should not fade with oxidation or temperature.

- Intensity of colour: colour should be intense enough to detect small amounts of constituents and for making accurate results in low concentrations.

- Clarity of solution: the substance to be observed should be completely soluble in the solution

- Specificity: color produced should be specific for the desired constituent

- Validity of Beer’s law: the intensity of the color should be proportional to the concentration which can be shown by plotting the Absorbance vs conc graph and a straight line passing through the origin should be obtained.

Some of the things that you can consider for taking care of a colorimeter are:

- Users should always remove the cuvette from the instrument when not using it.

- If any optical marks are found on the cuvette, gently clean it with tissue paper.

- Switch off the colorimeter when not in use. This will help the lamp to have a much more lasting life and will also save energy.

- After the work is done, it is preferred to disconnect the plug from the switchboard and cover the colorimeter with its’ protection cover.

- Make it a habit to always check the main power adapter and cable to know if there is any wear or tear. Prolonged use of the colorimeter instrument with even the slightest wear and tear can not only damage the machine but is also dangerous.

- Replace if the instrument is damaged.

- Always keep it in cool places and at room temperature.

- Do not keep it near harmful chemicals or burning materials.

Applications of a Colorimeter

- It is employed in the printing industries to evaluate the caliber of printing ink and paper.

- These are used in the food and food processing industries.

- It is frequently used in laboratories and hospitals to determine the biochemical composition of samples like blood, urine, cerebral spinal fluid, plasma, serum, etc.

- It is used in the textile and paint industries.

- Diamond dealers also examine the visual characteristics of priceless stones using a colorimeter.

- The instrument is also employed in cosmetology to measure the UV protection level of skin-care products.

- They are used to evaluate the water’s purity and screen for the presence of chemicals like cyanide, iron, fluorine, chlorine, molybdenum, etc.

- They are employed to evaluate the color contrast and brightness of screens on mobile devices, computers, and televisions to give people the greatest viewing experience.

- A colorimeter is also employed in the pharmaceutical sector to spot inferior goods and medications.

- Blood samples are tested using a colorimeter to determine the amount of hemoglobin present.